Non-invasive haemodynamic monitoring in Pregnancy: A comparative study using Ultrasound and Bioreactance

Authors: D Vinayagam, S Bowe, E Sheehan, B Thilaganathan, A Khalil

Fetal-Maternal Medicine Unit, St George’s University Hospitals NHS Foundation Trust, University of London

Keywords: Non-invasive cardiac output monitoring, Pregnancy, USCOM®, Bioreactance, NICOM®

Corresponding Author

Dr Asma Khalil

Consultant in Maternal Fetal Medicine

Fetal-Maternal Medicine Unit

Lanesborough Wing

St George’s University Hospitals NHS Foundation Trust

University of London

Blackshaw Road

London SW17 0QT

United Kingdom

Telephone: +44 20 8725 0071
Fax: +44 20 8725 0079    
E-mail: asmakhalil79@googlemail.com

ABSTRACT

Background

Due to the importance of cardiovascular dysfunction and advances in the development and use of non-invasive cardiac output monitoring devices, there is a growing interest in their use in the obstetric population. The aim of this study was to compare two commercially available, non-invasive cardiac output monitors in the assessment of heart rate (HR), stroke volume (SV), cardiac output (CO), cardiac index (CI) and total peripheral resistance (TPR) in an obstetric population.

Methods

This was a prospective, comparative study including two groups. A normotensive group consisted of uncomplicated healthy pregnancies; a hypertensive group consisted of those complicated by hypertensive disorders of pregnancy. Non-invasive cardiovascular assessments were peformed using two methodologies. USCOM® is an operator-dependent Doppler-ultrasound device which obtains velocity time integrals of transaortic blood flow at the left ventricular outflow tract, while NICOM® provides an operator-independent haemodynamic profile by assessing thoracic voltage changes (bioreactance) which reflect pulsatile blood flow in the thorax. Correlation coefficients were derived, and Bland-Altman analysis was performed to derive the mean percentage difference (MPD) between the devices.

Results

598 women were recruited to this study. In the normotensive group, 524 paired results were analysed, while 74 paired results were analysed in the hypertensive group. In the normotensive group, we found excellent correlation between USCOM® and NICOM® for HR (r= 0.885, p<0.05), and moderate correlations for SV (r=0.445, p<0.05), CO (r= 0.529, p<0.05), CI (r=0.385, p<0.05) and TPR (r=0.524, p<0.05). In the hypertensive group, we obtained similar correlations (HR r=0.877, p<0.05; SV r=0.575, p<0.05; CO r=0.601, p<0.05; TPR r=0.589, p<0.05). Bland Altman analysis found that the agreement between both methodologies improved as gestation advances, with a MPD of 34% for CO estimation in the third trimester of uncomplicated pregnancies, and 39% in pregnancies complicated by hypertensive disorders.

Conclusion

Our findings suggest that the two methodologies perform similarly in both uncomplicated pregnancies and in pregnancies complicated by hypertensive disorders. The study findings do not preclude the use of USCOM® and NICOM® devices in pregnancy, but indicate that these platforms cannot be used interchangeably. Our findings demonstrate an improvement in MPD as gestation advances, and therefore questions the validity of previous longitudinal studies investigating maternal hemodynamics using these methodologies. Our work also highlights the need to construct reference ranges for each device and for validation of each methodology against a reference method before their use in research or clinical practice.

Introduction

Pregnancy is associated with an array of haemodynamic changes which include a fall in the total peripheral resistance (TPR) and blood pressure (BP), and an increase in stroke volume (SV), cardiac output (CO) and heart rate (HR) (1). Aberrations in maternal hemodynamics have been reported to occur weeks prior to the onset of complications such as preeclampsia and fetal growth restriction (FGR). Whilst maternal mortality in the UK is decreasing sepsis, haemorrhage and preeclampsia are amongst the leading causes of direct maternal deaths (2). These life threatening illnesses require early and aggressive fluid resuscitation in order to limit tissue hypoxia and thereby prevent end organ damage, ultimately the cause of maternal mortality. Equally, excessive fluid resuscitation can have significant respiratory and neurological sequelae, such as with preeclampsia.

Despite significant advances in our knowledge of shock, our clinical management is based on traditionally measured variables including pulse, peripheral blood pressure, arterial oxygen saturation and central venous pressure. These act as surrogate markers for maternal SV and CO, which are fundamental to the appropriate management of the critically-ill obstetric patient. We know that in the early stages of critical illness, these proxy markers exhibit minimal change (3), and in a population consisting of largely young, medically fit women, decompensation in these surrogate variables is seen fairly late on in the disease process.

In previous decades, changes in maternal hemodynamics were investigated using dye dilution techniques and pulmonary artery catheterization (PAC) (4, 5). CO measurement with a PAC using the bolus thermodilution method has been considered as the gold standard (6, 7). Recently, there have been advances in the development and use of non-invasive CO monitoring devices. The technologies vary from ultrasound, carbon dioxide rebreathing, pulse contour analysis, bioimpedance and bioreactance. Two commercially available non-invasive devices are the USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non invasive cardiac output monitor, Cheetah Medical, USA). Comparative data for different modalities of non-invasive cardiac monitoring in the pregnant population is scarce. Validation data of non-invasive technology in pregnancy, against the gold standard of PAC is practically and ethically unobtainable. Our group has conducted a comparison study of ultrasound and bioreactance technology to transthoracic echocardiography and have demonstrated acceptable levels of agreement (manuscript currently undergoing peer review).

The aim of this study was to investigate the correlation and agreement between the NICOM® (thoracic bioreactance) and USCOM® (continuous wave Doppler ultrasound) for the estimation of heart rate (HR), stroke volume (SV), cardiac output (CO), cardiac index (CI) and total peripheral resistance (TPR) in both healthy pregnancies and pregnancies complicated by hypertensive disorders.

Methods

Study population and recruitment

This was a prospective, comparative study of 598 pregnant women. Women aged 16 and over were recruited from routine antenatal clinics at our tertiary centre. Women with a history of congenital or acquired heart disease were excluded. Women who had a diagnosis of a hypertensive disorder of pregnancy according to ISSHP diagnostic criteria (8), or were found to have a mean arterial pressure greater than 125mmHg at examination, were recruited to the hypertensive group. All subjects in the hypertensive group were in the third trimester. Written consent was obtained from all study participants and research ethics committee approval (12/LO/0810) was obtained prior to performing the study investigations.

All non-invasive assessments were performed in the same room, under standardised conditions for the entire cohort.  Subjects in the first and early second trimester (before 20 weeks’ gestation) were assessed in supine position, while those in the late second trimester (after 20 weeks’ gestation) and third trimester were in a left lateral position. The investigating team received formal training on the correct use of USCOM® and NICOM® by representatives of the manufacturers.

Research investigations

USCOM® employs continuous wave Doppler, with a non-imaging probe in the suprasternal notch to obtain velocity time integrals (VTI) of transaortic or transpulmonary blood flow at the left or right ventricular outflow tract respectively. Using an anthropometric algorithm, which correlates the outflow tract diameter with the patient’s given height, USCOM® uses the VTIs to compute SV and therefore CO and a complete hemodynamic profile. The USCOM® probe was placed in the suprasternal notch to obtain velocity time integrals (VTI) of transaortic blood flow at the left ventricular outflow tract. Each acquisition used for analysis had a minimum of 2 consecutive Doppler profiles. Acquisitions with the least amount of interference and the best quality VTIs were obtained in flowtracer mode.

The Non-Invasive Cardiac Output Monitor, NICOM® (Cheetah Medical, USA) uses thoracic bioreactance technology which is based on the analysis of thoracic voltage amplitude changes in response to a high-frequency injected current. The NICOM® device analyses the variations in frequency spectra (relative phase shifts) after delivering a transthoracic alternating current and uses several assumptions (about thoracic shape and fluid volumes) and algorithms, to compute SV, CO and an array of hemodynamic parameters. The phase shifts are measured continuously, and have been shown to correlate almost linearly to blood flow in the aorta (9). NICOM® surface electrodes were placed on participants (posterior aspect of the thorax) prior to them lying supine / left lateral position, and readings were obtained simultaneously to USCOM® readings. Following machine calibration, one complete haemodynamic profile was recorded for each patient.

Statistical analysis

The Shapiro-Wilk test was performed on all data sets to assess the distribution of data. The correlation coefficients were expressed as either Pearson’s (parametric data) or Spearman’s (non-parametric data). Categorical data were presented as number (%), while continuous data were presented as median and interquartile range (IQR). Bland-Altman analysis was performed, and the mean percentage difference (MPD) calculated as a ratio of two standard deviations of the bias (limit of agreement) to the mean value of the variable between the two methodologies. If the MPD is less than 30%, then  it is considered to be clinically acceptable, as proposed by Critchley and Critchley (6). All statistical analysis was carried out using IBM SPSS Statistics version 21.

Results

A total of 598 paired results were obtained between October 2013 and September 2015. The demographic details of the cohort are shown in Table 1. The median values with the interquartile range for each of the variables in both groups are shown in Table 2. In order to assess the linear relationship between the two methodologies, correlation coefficients were calculated for each variable for both groups (Table 3). Figures 1a-e illustrate the scatterplots providing a visual demonstration of the correlation between the two methodologies for HR, SV, CO, CI and TPR in the normotensive pregnancies. Figures 2a-e illustrate the correlations in the pregnancies complicated by hypertensive disorders of pregnancy. The mean percentage differences between the two methodologies are presented in Table 4.

Both devices showed high levels of correlation (normotensive group, r=0.885, p<0.05; hypertensive group r=0.877, p<0.05) and agreement (MPD of 15% and 16%) for the estimation of HR in healthy and complicated pregnancies. The correlations for stroke volume (normotensive group, r=0.445, p<0.05; hypertensive group r=0.575, p<0.05), cardiac output (normotensive group, r=0.529, p<0.05; hypertensive group, r=0.601, p<0.05), cardiac index (normotensive group, r=0.385, p<0.05; hypertensive group, r=0.312, p<0.05) and total peripheral resistance (normotensive group, r=0.524, p<0.05; hypertensive group r=0.589, p<0.05) were similar in both groups.

Bland Altman plots demonstrating the mean bias between the two methods and 95% limits of agreement, for each variable are shown for the normotensive group (Figures 3a-e) and the hypertensive group (Figures 4a-e). In the normotensive group the 95% limits of agreement (LOA) was -14.95bpm to 10.77bpm and the mean bias was -2.09bpm for HR, -32.77mls to 37.93mls and 2.58mls for SV, -2.74L/min to 2.83L/min and 0.04L/min for CO, -1.68L/min/m2 to 1.52L/min/m2 and -0.08L/min/m2 for CI and -771.0 dynes-sec-cm5 to 606.0 dynes-sec-cm5 and -83.0 dynes-sec-cm5 for TPR, respectively. In the pregnancies complicated by hypertensive disorders in pregnancy the 95% limits of agreement (LOA) was -16bpm to 12bpm, with a mean bias of -2bpm for HR, -31mls to 30mls and 0.2mls for SV, -2.96L/min to 2.54L/min and -0.2L/min for CO, -1.8L/min/m2 to 1.28L/min/m2 and -0.26L/min/m2 for CI and -688 dynes-sec-cm5 to 652 dynes-sec-cm5 and -18 dynes-sec-cm5 for TPR, respectively.

The MPD was <30% for HR estimation, while it was >30% for the estimation of SV, CO, CI and TPR. Analysis of the control group by trimester, demonstrates that the agreement between both methodologies improved as the pregnancy advances (Table 4), with an observed MPD of 34% for CO estimation in the third trimester of healthy pregnancies and MPD of 39% in pregnancies complicated by hypertensive disorders.

Both devices are able to deduce the CI (cardiac output corrected for body surface area). Due to differences in device algorithms, both devices do not derive the same body surface area (BSA), despite identical height and weight data being inputted in each device. NICOM® calculates the BSA according to the Du Bois formula, while USCOM® employs a different formula that gives a different BSA. Therefore, it is unsurprising that the CI has the weakest correlation (normotensive group, r=0.385, p<0.05; hypertensive group, r=0.312, p<0.05) of all the variables assessed. As a result, we did not analyse the stroke volume index (SVI) or total peripheral resistance index (TPRI) in either group.

Discussion

Summary of the study findings

Our results suggest that there is a moderate correlation between USCOM® and NICOM®, in recording SV, CO, CI and TPR, while that for HR was strong. Upon analysis of the data for each trimester, our findings suggest that the correlations for SV, CO, CI and TPR strengthen as the pregnancy advances. Similarly, the MPD revealed better agreement between the two methodologies in the third compared to the first or second trimesters. Importantly, we have demonstrated similar correlations, Bland Altman agreement values and MPD in both uncomplicated and pathological pregnancies.

Interpretation of the study findings and comparison with the existing literature

This is the first large, prospective study comparing two non-invasive methodologies in a large obstetric cohort consisting of both healthy and complicated pregnancies, ranging from 10-41 weeks of gestation. Our findings have shown that the two non-invasive methodologies perform similarly in normal and pathological pregnancies. This is an important finding as there would be little relevance if this technology performed differently in healthy compared to pathological pregnancies.

There is a paucity of data validating non-invasive technology in the pregnant population. Our group has conducted a comparison between non-invasive technology and transthoracic echocardiography, and we found acceptable levels of agreement between both methodologies and echocardiography in advanced pregnancy (manuscript currently undergoing peer review). Similar findings comparing USCOM to 3-d echocardiography in advanced pregnancy have been reported (10). The majority of studies in the non-pregnant population suggest that both USCOM® and NICOM® have a MPD exceeding the 30% threshold (11-14); this appears to be similar in pregnancy.

Bland Altman statistical analysis remains the most frequently used method of analysis when comparing two cardiac output monitoring devices. While a threshold of <30% MPD has been stated to be clinically acceptable, this is applicable only when one is assessing a new device against a reference method or gold standard. A MPD of <30% implies that the devices can be used interchangeably, or that a new method can replace the reference method or gold standard. However in the current study, we are comparing two new methodologies against each other, and therefore, achieving less than the 30% threshold in this context is likely to be clinically irrelevant.

Study limitations and strengths

The main strength of our study is the large sample size and the wide range of gestational age at cardiovascular assessment (10 to 41 weeks of gestation), allowing us to assess the correlation and agreement throughout pregnancy. A further strength of our study is that we have compared the two methodologies in both uncomplicated healthy and in pathological pregnancies, as ultimately, the primary research focus and clinical application of this technology will be in those complicated pregnancies. The main limitation of this study is that it compares two non-invasive methodologies against each other, without a reference method. Whilst we can assess the correlation and agreement between the two methods, it does not inform us of their accuracy against an accepted reference method.

Clinical and research implications

We have demonstrated that there is equal performance between the two methodologies in healthy pregnancies and in pregnancies complicated by hypertensive diseases. Our findings demonstrate that USCOM® and NICOM® cannot be used interchangeably in the pregnant population, due to the MPD being higher than 30%. We have demonstrated that the MPD for SV and CO estimation to be at least 50% in the first and second trimesters of pregnancy. This demonstrates an unacceptable level of variation for these platforms to be used interchangeably in a research or clinical setting. The variation in MPD with advancing gestation brings into question the validity and interpretation of longitudinal studies of maternal hemodynamics using these or similar platforms that have not been comprehensively validated in pregnancy. Further studies, ideally using a cross-sectional obstetric population and an accepted reference method, are required to ascertain whether the wide variation we have found in this study in the first two trimesters of pregnancy is attributable to the poor performance of one, or both platforms. Our work also highlights the need to devise specific reference ranges for each methodology (or device). We believe that further evaluation of these platforms is required before their use in research or the management of the critically unwell patient.

Conclusion

We have demonstrated excellent levels of correlation for HR, and moderate levels of correlation for SV, CO and TPR between Doppler ultrasound and bioreactance-derived haemodynamic indices in a large obstetric cohort. We found similar correlations in both healthy and pathological pregnancies, thereby suggesting that these devices can be utilised in normal pregnancies as well as those complicated by hypertensive disorders. However, clinically acceptable agreement between the two methods is lacking, and therefore they cannot be used interchangeably. Further validation studies and establishing device-specific reference ranges are needed before their application in research studies or the management of critically ill obstetric patients.

References

1.	Duvekot JJ, Peeters LL. Maternal cardiovascular hemodynamic adaptation to pregnancy. Obstetrical & gynecological survey. 1994;49(12):S1.

2.	Marian Knight SK, Peter Brocklehurst, Jim Neilson,, Judy Shakespeare JJK. Saving Lives, Improving Mothers Care Lessons learned to inform future maternity care from the UK and Ireland Confidential Enquiries into Maternal Deaths and Morbidity 2009-2012 2014.

3.	Marik PE, Baram M, Vahid B. Does central venous pressure predict fluid responsiveness?: a systematic review of the literature and the tale of seven mares. CHEST Journal. 2008;134(1):172-8.

4.	Ackerman WE, Porembka DT, Juneja MM, Kaczorowski DM. Use of a pulmonary artery catheter in the management of the severe preeclamptic patient. Anesthesiology review. 1991;17(4):37-40.

5.	Ueland K, Gills R, Hansen J, WYANT GM. Maternal cardiovascular dynamics. I. Cesarean section under subarachnoid block anesthesia. Survey of Anesthesiology. 1969;13(1):67-70.

6.	Critchley LH, Critchley JJH. A Meta-Analysis of Studies Using Bias and Precision Statistics to Compare Cardiac Output Measurement Techniques. Journal of Clinical Monitoring and Computing. 1999;15(2):85-91.

7.	Critchley LA, Lee A, Ho AM. A critical review of the ability of continuous cardiac output monitors to measure trends in cardiac output. Anesth Analg. 2010;111(5):1180-92.

8.	Tranquilli AL, Dekker G, Magee L, Roberts J, Sibai BM, Steyn W, et al. The classification, diagnosis and management of the hypertensive disorders of pregnancy: A revised statement from the ISSHP. Pregnancy Hypertens. 2014;4(2):97-104.

9.	Funk DJ, Moretti EW, Gan TJ. Minimally invasive cardiac output monitoring in the perioperative setting. Anesth Analg. 2009;108(3):887-97.

10.	McNamara H, Barclay P, Sharma V. Accuracy and precision of the ultrasound cardiac output monitor (USCOM 1A) in pregnancy: comparison with three-dimensional transthoracic echocardiography. Br J Anaesth. 2014;113(4):669-76.

11.	Chong SW, Peyton PJ. A meta-analysis of the accuracy and precision of the ultrasonic cardiac output monitor (USCOM). Anaesthesia. 2012;67(11):1266-71.

12.	Knirsch W, Kretschmar O, Tomaske M, Stutz K, Nagdyman N, Balmer C, et al. Cardiac output measurement in children: comparison of the Ultrasound Cardiac Output Monitor with thermodilution cardiac output measurement. Intensive Care Medicine. 2008;34(6):1060-4.

13.	Kupersztych-Hagege E, Teboul JL, Artigas A, Talbot A, Sabatier C, Richard C, et al. Bioreactance is not reliable for estimating cardiac output and the effects of passive leg raising in critically ill patients. Br J Anaesth. 2013;111(6):961-6.

14.	Tan H, Pinder M, Parsons R, Roberts B, Van Heerden P. Clinical evaluation of USCOM ultrasonic cardiac output monitor in cardiac surgical patients in intensive care unit. British Journal of Anaesthesia. 2005;94(3):287-91.

Table 1 Baseline characteristics of the entire cohort consisting of the normotensive (n=524) and hypertensive groups (n=74).

|  | First Trimester (n=128) | Second Trimester (n=90) | Third Trimester (n=306) | Normotensive Group (n=524) | Hypertensive Group (n=74) |
| --- | --- | --- | --- | --- | --- |
| Maternal age in years, median (IQR) | 31 (8.4) | 32 (5.6) | 31 (8.0) | 32 (8.2) | 31 (8.2) |
| Nulliparity, n (%) | 68 (53) | 53 (59) | 165 (54) | 286 (55) | 42 (57) |
| Caucasian, n (%) | 52 (41) | 38 (42) | 183 (59) | 273 (52) | 45 (61) |
| Maternal weight at booking (kg), median (IQR) | 65.1 (15.6) | 68.7 (22.1) | 66.3 (20.5) | 66.0 (18.7) | 72.7 (23.9) |
| Maternal mean arterial pressure at booking (mmHg), median (IQR) | 82 (12.6) | 83.9 (12.7) | 83.1 (13) | 83.3 (12.6) | 92 (14) |
| Maternal weight at assessment (kg), median (IQR) | 65.4 (14.7) | 72.2 (22.4) | 77.9 (20.9) | 73.2 (20.3) | 83.5 (20.2) |
| Maternal mean arterial pressure at assessment (mmHg), median (IQR) | 83.6 (10.6) | 82.4 (17.4) | 90.1 (0.9) | 87.9 (10.5) | 106 (11.6) |

Table 2. Median values with the interquartile range (IQR) of the hemodynamic variables for the cohort obtained by USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non invasive cardiac output monitor, Cheetah Medical, USA) in both groups.

|  | Heart Rate (beats/min) | Stroke Volume (ml) | Cardiac Output (L/min) | Cardiac Index (L/min/m2) | Total Peripheral Resistance (dyn. s. cm−5) |
| --- | --- | --- | --- | --- | --- |
| USCOM® |  |  |  |  |  |
| First Trimester | 70 (15) | 84 (18) | 5.93 (1.4) | 3.39 (0.84) | 1143 (306) |
| Second Trimester | 79 (12) | 83 (23) | 6.59 (1.65) | 3.47 (0.87) | 1037 (253) |
| Third Trimester | 87 (19) | 76 (20) | 6.60 (1.45) | 3.41 (0.81) | 1080 (321) |
| Normotensive Group | 81 (20) | 79 (21) | 6.41 (1.57) | 3.41 (0.82) | 1090 (313) |
| Hypertensive Group | 85 (20) | 82 (14) | 6.88 (1.96) | 3.29 (0.75) | 1293 (372) |
| NICOM® |  |  |  |  |  |
| First Trimester | 72 (15) | 74 (26) | 5.60 (2.45) | 3.3 (1.3) | 1161 (541) |
| Second Trimester | 82 (13) | 77 (24) | 6.03 (2.55) | 3.4 (1.1) | 1114 (407) |
| Third Trimester | 89 (17) | 80 (23) | 6.95 (1.87) | 3.77 (0.83) | 1064 (334) |
| Normotensive Group | 83 (19) | 78 (25) | 6.46 (2.03) | 3.6 (0.96) | 1092 (395) |
| Hypertensive Group | 86 (20) | 81 (22) | 7.05 (2.42) | 3.71 (0.93) | 1295 (495) |

Table 3 Correlation coefficients (Pearson’s / Spearman’s) between USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non-invasive cardiac output monitor, Cheetah Medical, USA) for each variable. Correlation coefficients for each trimester of, and the entire normotensive group are presented, along with the correlation coefficients in the hypertensive group consisting of pregnancies complicated by hypertensive disorders of pregnancy.

|  | Heart Rate (beats/min) | Stroke Volume (ml) | Cardiac Output (L/min) | Cardiac Index (L/min/m2) | Total Peripheral Resistance (dyn. s. cm−5) |
| --- | --- | --- | --- | --- | --- |
| First Trimester | 0.891 | 0.395 | 0.389 | 0.336 | 0.446 |
| Second Trimester | 0.831 | 0.385 | 0.431 | 0.232 | 0.393 |
| Third Trimester | 0.850 | 0.598 | 0.553 | 0.438 | 0.602 |
| Normotensive Group | 0.885 | 0.445 | 0.529 | 0.385 | 0.524 |
| Hypertensive Group | 0.877 | 0.575 | 0.601 | 0.312 | 0.589 |

Table 4 Mean percentage difference (MPD) (%) between USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non-invasive cardiac output monitor, Cheetah Medical, USA) for each variable. The MPD is the ratio of the limit of agreement (2 standard deviations of the bias between the two methods) and the mean value obtained by both methodologies. A MPD of <30% indicates that there is sufficient agreement between two methodologies so that they may be used interchangeably, or so that one method can be replaced by another. Mean percentage differences for each trimester of, and the entire normotensive group are presented, along with the MPD we found in the hypertensive group consisting of pregnancies complicated by hypertensive disorders of pregnancy.

|  | Heart Rate (beats/min) | Stroke Volume (ml) | Cardiac Output (L/min) | Cardiac Index (L/min/m2) | Total Peripheral Resistance (dyn. s. cm−5) |
| --- | --- | --- | --- | --- | --- |
| First Trimester | 13 | 52 | 54 | 55 | 77 |
| Second Trimester | 13 | 52 | 49 | 57 | 65 |
| Third Trimester | 16 | 35 | 34 | 37 | 44 |
| Normotensive Group | 15 | 45 | 43 | 46 | 60 |
| Hypertensive Group | 16 | 38 | 39 | 44 | 50 |

Figure Legends

Figure 1. Scatterplot of heart rate (1a), stroke volume (1b), cardiac output (1c), cardiac index (1d) and total peripheral resistance (1e) estimations by USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non-invasive cardiac output monitor, Cheetah Medical, USA) in normotensive pregnancies. The reference lines represent a correlation coefficient of 1.00 for comparison.

Figure 2. Scatterplot of heart rate (2a), stroke volume (2b), cardiac output (2c), cardiac index (2d) and total peripheral resistance (2e) estimations by USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non-invasive cardiac output monitor, Cheetah Medical, USA) in the pregnancies complicated by hypertensive disorders. The reference lines represent a correlation coefficient of 1.00 for comparison.

Figure 3. Bland Altman Plot of the heart rate (3a), stroke volume (3b), cardiac output (3c), cardiac index (3d) and total peripheral resistance (3e) estimations by USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non invasive cardiac output monitor, Cheetah Medical, USA) in normotensive pregnancies.

Figure 4. Bland Altman Plot of the heart rate (4a), stroke volume (4b), cardiac output (4c), cardiac index (4d) and total peripheral resistance (4e) estimations by USCOM® (ultrasound cardiac output monitor, USCOM Ltd, NSW, Australia) and NICOM® (Non invasive cardiac output monitor, Cheetah Medical, USA) in the pregnancies complicated by hypertensive disorders.