Analytical Evaluation of the automated Galectin-3 assay on the Abbott ARCHITECT immunoassay instruments

David C. Gaze1, Christian Prante2, Jens Dreier2, Cornelius Knabbe2, Corinne Collet3, Jean-Marie  Launay3, Janka Franekova4,  Antonin Jabor4, Lieselotte Lennartz5, Jessie Shih6, Jose Manuel del Rey7, Martina Zaninotto8, Mario Plebani8, Paul O. Collinson1

1Chemical Pathology, Clinical Blood Sciences, St George’s Healthcare NHS Trust, London, United Kingdom

2Heart and Diabetes Center North Rhine Westfalia, Institute for Laboratory and Transfusion Medicine, Bad Oeynhausen, Germany

3Laboratory of Clinical Biochemistry, Lariboisière Hospital, Paris, France

4Department of Laboratory Methods, Institute for Clinical and Experimental Medicine, Prague, Czech Republic

5 Abbott Laboratories, Wiesbaden, Germany

6 Abbott Laboratories, Abbott Park, IL, US

7Clinical Chemistry, Core Laboratory, Ramon y Cajal Hospital, Madrid, Spain

8Department of Laboratory Medicine, University Hospital of Padova, Italy

- Address for Correspondence:

- David C. Gaze

Chemical Pathology, Clinical Blood Sciences,

St George’s Healthcare NHS Trust,

Blackshaw Road

London, SW17 0QT

United Kingdom

Tel:	 +44 (0) 208 725 5878

Fax:	+44 (0) 208 725 5868

Email: 	david.gaze@stgeorges.nhs.uk

Running title: Evaluation of ARCHITECT Galectin-3 Immunoassay

Key Words: Galectin-3, Chronic heart failure, Analytical Performance, ARCHITECT

Non Standard Abbreviations: CHF, chronic heart failure; CLSI, Clinical Laboratory Standards Institute; %CV, coefficient of variation; LoB, Limit of Blank; LoD, Limit of Detection; LoQ, Limit of Quantification; AMR, analytical measurement range.

Acknowledgments: Abbott Laboratories provided the protocol and reagents for this study.

- Number of Words: 2710

- Number of Figures: 3

- Number of Tables: 3

- Number of References: 22

- Abstract

Aim. Galectin-3 is secreted from macrophages and binds and activates fibroblasts forming collagen. Tissue fibrosis is central to the progression of chronic heart failure (CHF). We performed a European multi-centred evaluation of the analytical performance of the two-step routine and Short Turn-Around-Time (STAT) Galectin-3 immunoassay on the ARCHITECT i1000SR, i2000SR and i4000SR (Abbott Laboratories).

Materials and methods. We evaluated the assay precision and dilution linearity for both routine and STAT assays and compared serum and plasma and fresh versus frozen samples. The reference interval and biological variability were also assessed. Measurable samples were compared between ARCHITECT instruments and between the routine and STAT assays and also to a Galectin-3 ELISA (BG Medicine Inc.).

Results. The total assay coefficient of variation (CV%) was 2.3-6.2% and 1.7 to 7.4% for the routine and STAT assays respectively. Both assays demonstrated linearity up to 120 ng/mL. Galectin-3 concentrations were higher in plasma samples than in serum samples and correlated well between fresh and frozen samples (R=0.997), between the routine and stat assays, between the ARCHITECT i1000 and i2000 instruments and with the Galectin-3 ELISA. The reference interval on 627 apparently healthy individuals (53% male) yielded upper 95th and 97.5th percentiles of 25.2 ng/mL and 28.4 ng/mL respectively. Values were significantly lower in subjects aged <50 years.

Conclusions. The Galectin-3 Routine and STAT assays on the Abbott ARCHITECT instruments demonstrated good analytical performance. Further clinical studies are required to demonstrate the diagnostic and prognostic potential of this novel marker in patients with CHF.

- Introduction

Heart failure is common in Europe, affecting approximately 1-2% of the adult population (1). The incidence and prevalence increases with age with the average age of diagnosis being 76 years (2). The risk of CHF is greater in men than women but there are more women with CHF than men due to population demographics (1). Prognosis is poor; forty percent of patients die within one year of diagnosis and survival is less than 10% per year thereafter (3). Echocardiography and biochemical assays for natriuretic peptides (NTproBNP or BNP) are routinely used for diagnosis and management of cardiac heart failure patients (1).

Galectin-3 (30kDa, LGALS3, Chr 14,q21-q22) is a -galactoside-binding lectin secreted from macrophages (4). Galectin-3 is involved in inflammation and in the pathogenesis of tumour growth and autoimmune conditions (5). Galectin-3 binds to and activates fibroblasts forming collagen and results in fibrosis (scaring) (6). Galectin-3 has been associated with tissue repair, cardiac inflammation and cardiac remodelling (6-11). Tissue fibrosis is the final common pathway central to the progression of CHF. Diffuse cardiac fibrosis impairs ventricular function and contributes to both systolic and diastolic dysfunction. Galectin-3 therefore is a potential diagnostic and prognostic marker of CHF as well as a surrogate marker to assess therapeutic intervention aimed at reducing the progression of CHF (12-15).

The enzyme-linked immunosorbent assay (ELISA) commercially available for the measurement of Galectin-3 presents satisfactory quality specifications but requires some manual steps (16). High throughput assays allowing automation and a shorter turnaround time on immunoassay analysers will be required for measurement of Galectin-3 in routine clinical use. We performed a multi-centred evaluation of the analytical performance of the two-step routine and Short Turn-Around-Time (STAT) Galectin-3 immunoassays on the ARCHITECT i1000SR, i2000SR and i4000SR (Abbott Laboratories, Wiesbaden, Germany) at six European sites.

- Materials and Methods:

- Sites, Instruments and Reagents

The ARCHITECT Galectin-3 assay is a chemiluminescent microparticle immunoassay for the quantitative determination of Galectin-3 in human serum or EDTA plasma performed on the ARCHITECT i systems.  The assay is a two-step protocol. In step 1, 25 L of sample and M3/M38 anti-Galectin-3 coated paramagnetic microparticles are incubated. Galectin-3 molecules present in the sample bind the M3/M38 antibody. Following a wash step the second step incorporates the addition of a B7/B5 anti-Galectin-3 acridinium-labelled conjugate, creating an antibody-antigen sandwich. After a further washing cycle, a pre-trigger and trigger solution is added to the reaction vessel creating a chemiluminescent reaction. The light emission is measured as relative light units by a photomultiplier and is directly proportional to the amount of Galectin-3 in the sample. The routine assay takes 29 minutes. The short turn around time (STAT) assay also uses 25 L of sample and can be processed in 18 minutes.

The routine assay was assessed in the UK on an i2000SR (site 1) and in Germany on an i4000SR (site 2). The STAT version of the assay was also investigated in the UK (i2000SR) as well as in France (site 3, on two i2000SR instruments), the Czech Republic (site 4), Italy (site 6), on the i2000SR and in Spain (site 5) on the i1000SR.

An ELISA plate assay (BG Medicine, Inc. Waltham MA, USA) was available for methodological comparison.  The limit of blank (LoB) was 0.86 ng/mL; limit of Detection (LoD) was 1.13 ng/mL and the limit of Quantification (LoQ) was 1.32 ng/mL. The coefficient of variation %CV was 4.2 to 12.0% at 6.1 to 72.2 ng/mL. The dynamic measuring range of the ELISA assay was 1.32 to 96.6 ng/mL.

- Assay Precision and Limits of Blank, Detection and Quantification

The assay precision was determined at each site following the Clinical and Laboratory Standards Institute (CLSI) EP5-A2, based on 20 replicates of the low, medium and high concentration assay quality controls. QC materials were assayed in duplicate in two runs, separated by a minimum of two hours and run over five days.

The LoB, LoD and LoQ were verified for both the ARCHITECT Galectin-3 routine and STAT assays and performed according to CLSI guideline EP17-A on each analyzer. For LoB, 20 replicates of a blank material (Calibrator A, 0.0 ng/mL) were tested. For LoD verification, 20 replicates of a sample with a concentration close to the manufacturers claimed LoD were used. LoQ was verified by testing a minimum of 30 replicates of a sample with a concentration near the claimed LoQ. The LoB, LoD and LoQ ranged from 0.2-0.3, 0.6-0.7 and 5.0-<6.0 ng/mL respectively for the routine assay and 0.0-1.2, 0.5-1.7 and 2.1-4.0 ng/mL for the STAT assay, respectively (Table 1).

- Dilution Linearity

Linearity was assessed by making 5 dilutions of a high patient sample containing an analyte concentration near the upper end of the analytical measurement range (AMR), with a low patient sample.  Each dilution was tested in duplicate. Both the routine and STAT assays demonstrated dilution linearity up to 120 ng/mL (regression equation -2.454 + 1.036x).

- Sample matrix Comparison

Matched serum and EDTA plasma samples were drawn from 54 subjects and assayed fresh on the Galectin-3 STAT assay.

- Fresh Samples compared to Frozen Samples

The effects of sample freezing were assessed. EDTA plasma samples (n=36) were collected in the morning and stored at 2-8 oC. Samples were divided into two aliquots. 1 aliquot was tested on the same day for both methods, the other aliquot was frozen at -30oC for 2 days, thawed, mixed and reassayed on both methods.

- Methodological Comparison

Samples with concentrations of Galectin-3 across the measuring range were compared in three experiments. Firstly, comparisons were made between the routine and STAT assay. Secondly, comparisons were made between the i1000SR and i2000SR instruments and thirdly, samples were compared between the ARCHITECT Galectin-3 and the BG Medicine Galectin-3 ELISA.

Reference Interval

The reference interval was investigated using 627 samples from apparently healthy individuals with no evidence of cardiac disease. Site 1 (n=120) used fresh serum from apparently cardiac healthy individuals with negative serum cTnI as defined by the Siemens Centaur TnI-Ultra assay and normal NTproBNP concentrations determined on the Siemens IMMULITE 2000 XPi. Site 2 (n=130) used frozen plasma and site 6 (n=136) used fresh serum samples from unscreened apparently healthy volunteers. Site 3 (n=119) used frozen serum samples from volunteers self reporting the absence of a history of diabetes, hypertension, cardiovascular disease or cancer.  Site 5 (n=122) used fresh plasma samples from healthy individuals attending a routine medical check-up appointment. At all sites, the subject’s age and gender were recorded.

- Biological Variability

The Biological variability of Galectin-3 was determined in healthy individuals recruited at sites 2 (n=12) and 3 (n=18).  All subjects were bled twice within the same day, in the morning and the afternoon. In addition subjects recruited at site 2 were also bled for 5 consecutive days. Samples were assayed fresh for Galectin-3.

- Samples source, data handling and Statistics

Ethics approval was received for specimens collected for the reference interval and biological variation studies and waiver received for the use of surplus specimens. Data from each site were recorded in Microsoft Excel spreadsheets (Microsoft Corporation, Redmond WA, USA). Assay precision, linearity, reference intervals, Passing-Bablok and linear regression analyses were generated using Analyse-it (Analyse-it Software, Leeds, UK; http://analyse-it.com). The short term (within day) and long term (between 5 days) biological variability were calculated according to the methods of Fraser (17).

- Results

Imprecision based on manufacturer’s QC material was favourable and demonstrated total assay %CV of <7% for the routine assay and <8% for the STAT assay irrespective of platform. The imprecision for each site is summarised in Table 1. The results were similar between sites. The least precise QC occurred in the STAT assay using the low QC (9.3 ng/mL, %CV= 7.4%) whilst the most precise occurred in the high QC, (73.7 ng/mL, %CV= 1.7%) for the STAT assay.

Verification of the manufacturer’s claimed LoB, LoD and LoQ were determined by CLSI guidelines. The LoB, LoD and LoQ ranged from 0.2-0.3, 0.6-0.7 and 5.0-<6.0 ng/mL respectively for the routine assay and 0.0-1.2, 0.5-1.7 and 2.1-4.0 ng/mL for the STAT assay respectively. (Table1). Both the routine and STAT assays demonstrated dilutional linearity up to 120 ng/mL ng/mL (regression equation -2.454 +1.036x) (Supplemental material, Table 1 and Figure 1).

Observable differences between serum and plasma were assessed. Matched serum and plasma (n=54) samples taken at the same draw were measured for Galectin-3. There was no significant difference between serum and plasma in 10 samples measured at site 5 however using a larger population of 44 samples at site 6, Galectin-3 concentration were significantly higher in plasma compared to serum (p=<0.001). When combining the data the mean plasma Galectin-3 concentration was 39.5 ng/mL but was 36.8 ng/mL in serum samples (Passing-Bablok regression -1.60 +0.99x, Figure 1A; bias of -10.2%, 95%CI -36.8% to 16.3%, Figure 1B).

Galectin-3 concentrations were similar (r=0.997, Passing-Bablok regression 0.86 +1.07x, bias 10.1%, 95%CI -6.0% to 26.2%) between fresh samples (n=36, median Galectin-3 17.7 ng/mL) and frozen samples thawed once and reassayed (median Galectin-3 18.6 ng/mL). This was similar for the BG Medicine ELISA assay (r=0.999, Passing-Bablok regression 0.38 +1.11x, bias 12.8% (95%CI -3.5% to 29.0%) (Supplemental material, Table 2).

There was good correlation between the routine and STAT Galectin-3 assay formats (n=105, Passing-Bablok regression -0.52 +1.04x, bias 2.0%, 95%CI -11.0% to 15.1%, Figure 2A). Excellent correlation was observed between the Architect i1000 and i2000 instruments (n=30, Passing-Bablok regression 0.33 +0.99x, bias -0.8%, 95%CI -9.4-7.9%, Figure 2B). Consolidating the results from all sites in the largest comparison (n=807), there was good agreement between the Architect assay and the BG Medicine Galectin-3 ELISA (Passing-Bablok regression  -0.07 +1.04x, bias 6.7%, 95%CI -32.3% to 45.6%, Figure 2C).

A population of 627 apparently healthy individuals comprising 335 males and 292 females with no evidence of cardiac disease was used to determine the galectin-3 reference intervals. The combined data are summarised in Table 2 and presented as the upper 75th, 95th and 97.5th percentiles for the entire cohort and according to gender and age. The frequency distribution of Galectin-3 concentrations is shown in Figure 3A. The greatest frequency of concentrations occurred at approximately 12.5 ng/mL. A weak trend between increasing age and Galectin-3 concentration was observed (Figure 3B). Similar trends were observed in both males (Supplemental material, Figure 2A) and females (Supplemental material, Figure 2B).

The biological variability of Galectin-3 was assessed at two sites. The within subject (CVi) and between subject (CVg) biological variability were 10.3% and 22.4% respectively with a reference change value (RCV) of 30.7 for within day variation and 13.1% and 26.8% respectively with an RCV of 38.0 for between 5-day variability. Combining all available data from the two sites the total biological variability of Galectin-3 was 14.4% CVi and 23.3% CVg with an RCV of 43.5 (Table 3). There was a demonstrable difference in the variability of Galectin-3 in females compared to males (Supplemental material, Figure 3).

- Discussion

The performance characteristics of the ARCHITECT Galectin-3 automated immunoassay demonstrates that it is suitable both as a routine and short turn around time assay for the determination of Galectin-3 using the validated sample matrix, either serum or EDTA plasma. The assay demonstrated better precision than that reported previously for the BG Medicine ELISA assay (16). These data presented in the present multicentre European study are similar to precision data obtained in a two-site American study (18) that also obtained total %CV <6% across two reagent lots, when employing both the i1000SR and i2000SR instruments. The assay has a wide dynamic linear range suitable for the clinical utility in CHF patients. The assay is robustly stable with no significant difference between fresh and frozen samples. Similar findings of Galectin-3 stability have been observed over the course of two weeks when samples are stored either at room temperature or refrigerated or subjected to 6 freeze-thaw cycles (16).  In the present study Galectin-3 was higher in plasma than in serum (p=<0.001), however there is equal variance for the values in both matrices. This is contrary to the findings of La’ulu and colleagues (18) who demonstrated similar values between serum and EDTA plasma however they only compared 11 paired samples whereas the present study included 54. Additionally, Christenson et al (16) found similar values in 50 matched serum and EDTA plasma samples when using the BG Medicine ELISA assay.

A number of method comparisons were performed in the present study. There was good overall agreement between the two assay formats, between different instruments of the ARCHITECT family and between the automated assay and the comparator BG Medicine ELISA assay. Comparing 807 samples between the ARCHITECT and the comparator method, there was very little difference with a small positive bias of 6.7% towards the ARCHITECT method. These data suggest that clinical studies performed on the ELISA assay can be extrapolated to the ARCHITECT but the automated assay should be robustly investigated independently for its clinical utility.

The distribution of Galectin-3 concentrations in 1092 apparently healthy individuals has been studied previously using the BG Medicine ELISA assay. The 95th and 97.5th percentiles were found to be 20.3 ng/mL and 22.1 ng/mL respectively. The upper reference limits derived in the present study based on 627 individuals considered to be healthy were slightly higher at 25.2 and 28.4 ng/mL respectively for the automated ARCHITECT assay. In the comparative study from the US these values were much lower at <20 ng/mL for both the 95th and 97.5th percentile. It is probable that the smaller dataset (n=364) and a wider age distribution (range 19 to 79 years) in the US study is responsible for the disparity. De Boer and colleagues have previously investigated 7968 general population subjects as part of the Prevention of REnal and Vascular ENd-stage Disease (PREVEND) cohort and found a strong association between Galectin-3 and increasing age and gender (19). Concentrations of Galectin-3, determined using the BG Medicine ELISA, were greater in women than in men across the age range 30-75 years. A similar but less striking trend was observed in the current dataset (Table 2) which is probably attributable to the small gender specific populations studied. Larger datasets should be investigated using the automated ARCHITECT assay. A significant difference between male and female Galectin-3 concentrations was observed at sites 2 and 3 whist all other sites did not demonstrate a difference between the gender groups. No significant differences were demonstrated in the US study (18). Instrumentation and assay differences can be ruled out as an explanation as the sites differed by using i4000SR routine and i2000SR STAT Galectin-3 assays respectively. The large general population studied by de Boer showed divergence of Galectin-3 trend lines between males and females above the age of 50 years yet values were similar between the gender groups below 50 years old. It is possible that this phenomenon is responsible for lack of difference observed between the gender groups at most of the sites in this study. A further limitation to our reference interval study was the lack of ethnographic data, history of cardiovascular disease including diabetes mellitus or the use of cardiovascular medication. Furthermore we did not report renal function in the apparently healthy population. It has been established recently that elevated plasma Galectin-3 concentrations are associated with decreased estimated glomerular filtration rate (eGFR). O’Seaghdha and colleagues measured plasma Galectin-3 in 2,450 Framingham offspring participants. Elevated plasma Galectin-3 was associated with a rapid decline in eGFR (odds ratio 1.49, 95%CI 1.28-1.73) and to a higher risk of chronic kidney disease (odds ratio 1.47, 95%CI 1.27-1.71). Further studies are required to assess differences in Galectin-3 within the healthy population based on ethnicity, prior history of cardiovascular disease and associated risk factors including renal function status.

The use of age and gender specific reference limits has been proposed for other cardiac biomarkers such as NTproBNP, BNP and cardiac troponin rather than a single value and it is likely that the same should be applied to Galectin-3.

We report here to the best of our knowledge for the first time, the biological variability of Galectin-3 in the general population as determined by the ARCHITECT automated immunoassay. The biological variation of Galectin-3 based on the BG Medicine ELISA assay has been reported in a small study (n=17) previously (21).   The RCV’s reported in the present study are lower than the hourly RCV (39%) or the long term (8 weeks) RCV (61%) as previously reported. In addition the biological variation of Galectin-3 is lower than the long term biological variation of BNP (RCV 140%) which suggests Galectin-3 may be a better marker to guide therapy than BNP. Galectin-3 is a specific biomarker of fibrosis, a known indicator of CHF progression. Whilst the natriuretic peptides can also detect the consequences of fibrosis, they are influenced by fluid and volume overload (22) and thus may be less specific.

In summary this multi-centred European study has evaluated the analytical performance of the ARCHITECT Galectin-3 immunoassay. The assay demonstrated good imprecision and linearity and acceptable comparison to an ELISA method. An automated immunoassay is less labour intensive and facilitates faster turn around times than a manual ELISA, thus facilitating the execution of more clinical studies to further establish the relevance and application of Galectin-3 testing in clinical practice.

References

1.  McMurray JJV, Adamopoulos S, Anker SD, Auricchio A, Bohm M, Dickstein K et al.  ESC Guidelines for the diagnosis and treatment of acute and chronic heart failure 2012.  Eur Heart J. 2012;33:1787-1847.

2.  	Cowie MR, Wood DA, Coats AJ, Thompson SG, Poole-Wilson PA, Suresh V et al. Incidence and aetiology of heart failure; a population-based study. Eur Heart J 1999; 20(6):421-428.

3. Cowie MR, Wood DA, Coats AJ, Thompson SG, Suresh V, Poole-Wilson PA et al. Survival of patients with a new diagnosis of heart failure: a population based study. Heart 2000; 83(5):505-510.

4. Dumic J, Dabelic S, Flogel M. Galectin-3: an open-ended story. Biochim Biophys Acta 2006; 1760(4):616-635.

5. Henderson NC, Sethi T. The regulation of inflammation by galectin-3. Immunol Rev 2009; 230(1):160-171.

6. Henderson NC, MacKinnon AC, Farnworth SL, Poirier F, Russo FP, Iredale JP et al. Galectin-3 regulates myofibroblast activation and hepatic fibrosis. Proc Natl Acad Sci U S A 2006; 103(13):5060-5065.

7. 	Kuwabara I, Liu FT. Galectin-3 promotes adhesion of human neutrophils to laminin. J Immunol 1996; 156(10):3939-3944.

8. 	Liu YH, D'Ambrosio M, Liao TD, Peng H, Rhaleb NE, Sharma U et al. N-acetyl-seryl-aspartyl-lysyl-proline prevents cardiac remodeling and dysfunction induced by galectin-3, a mammalian adhesion/growth-regulatory lectin. Am J Physiol Heart Circ Physiol 2009; 296(2):H404-H412.

9. 	Papaspyridonos M, McNeill E, de Bono JP, Smith A, Burnand KG, Channon KM et al. Galectin-3 is an amplifier of inflammation in atherosclerotic plaque progression through macrophage activation and monocyte chemoattraction. Arterioscler Thromb Vasc Biol 2008; 28(3):433-440.

10. 	Sano H, Hsu DK, Yu L, Apgar JR, Kuwabara I, Yamanaka T et al. Human galectin-3 is a novel chemoattractant for monocytes and macrophages. J Immunol 2000; 165(4):2156-2164.

11. 	Sharma UC, Pokharel S, van Brakel TJ, van Berlo JH, Cleutjens JP, Schroen B et al. Galectin-3 marks activated macrophages in failure-prone hypertrophied hearts and contributes to cardiac dysfunction.   Circulation. 2004;110:3121-3128.

12. 	van Kimmenade RR, Januzzi JL, Jr., Ellinor PT, Sharma UC, Bakker JA, Low AF et al. Utility of amino-terminal pro-brain natriuretic peptide, galectin-3, and apelin for the evaluation of patients with acute heart failure. J Am Coll Cardiol 2006; 48(6):1217-1224.

13. 	Lok DJ, van der MP, de la Porte PW, Lipsic E, van Wijngaarden J, Hillege HL et al. Prognostic value of galectin-3, a novel marker of fibrosis, in patients with chronic heart failure: data from the DEAL-HF study. Clin Res Cardiol 2010; 99(5):323-328.

14. 	de Boer RA, Lok DJ, Jaarsma T, van der MP, Voors AA, Hillege HL et al. Predictive value of plasma galectin-3 levels in heart failure with reduced and preserved ejection fraction. Ann Med 2011; 43(1):60-68.

15. 	McCullough PA, Olobatoke A, Vanhecke TE. Galectin-3: a novel blood test for the evaluation and management of patients with heart failure. Rev Cardiovasc Med 2011; 12(4):200-210.

16. 	Fraser CG, Harris EK. Generation and application of data on biological variation in clinical chemistry. Crit Rev Clin Lab Sci 1989; 27(5):409-437.

17. 	Christenson RH, Duh SH, Wu AH, Smith A, Abel G, deFilippi CR et al. Multi-center determination of galectin-3 assay performance characteristics: Anatomy of a novel assay for use in heart failure. Clin Biochem 2010; 43(7-8):683-690.

18. 	La'ulu SL, Apple FS, Murakami MM, Ler R, Roberts WL, Straseski JA. Performance characteristics of the ARCHITECT Galectin-3 assay. Clin Biochem 2013; 46(1-2):119-122.

19. 	de Boer RA, van Veldhuisen DJ, Gansevoort RT, Muller Kobold AC, van Gilst WH, Hillege HL et al. The fibrosis marker galectin-3 and outcome in the general population. J Intern Med 2012; 272(1):55-64.

20. O’Seaghdha CM, Hwang SJ, Ho JE, Vasan RS, Levy D, Fox CS. Elevated Galectin-3 Precedes the Development of CKD. J. Am. Soc. Nephrol. 2013; 24:1470-1477.

21. 	Wu AH, Wians F, Jaffe A. Biological variation of galectin-3 and soluble ST2 for chronic heart failure: implication on interpretation of test results. Am Heart J 2013; 165(6):995-999.

22. 	Chen HH, Burnett JC. Natriuretic peptides in the pathophysiology of congestive heart failure. Curr Cardiol Rep 2000; 2(3):198-205.

Table 1: Total Assay precision (%CV) using quality control material (QC) and limits of Blank (LoB), Detection (LoD) and Quantification (LoQ).

|  | Site 1 |  | Site 2 | Site 3 |  | Site 4 | Site 5 | Site 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | UK | UK | Germany | France1 | France 2 | Czech R | Spain | Italy |
| Instrument | i 2000SR | i 2000SR | i 4000SR | i 2000SR | i 2000SR | i 2000SR | i 1000SR | i 2000SR |
| Assay format | Routine | STAT | Routine | STAT | STAT | STAT | STAT | STAT |
| Low QC (9.1 ng/mL) Mean +/- SD | 9.9 ± 0.6 | 9.8 ± 0.6 | 9.3 ± 0.6 | 9.9 ± 0.5 | 9.3 ± 0.7 | 8.8 ± 0.5 | 9.0 ± 0.4 | 9.3 ±0.4 |
| Low QC %CV | 6.2 | 5.9 | 6.0 | 4.5 | 7.4 | 5.3 | 4.6 | 3.8 |
| Medium (20.5 ng/mL) Mean +/- SD | 20.8 ± 0.6 | 20.8 ± 0.4 | 20.2 ± 0.9 | 21.2 ± 0.9 | 19.9 ± 0.7 | 19.5 ± 0.7 | 20.1 ± 0.5 | 19.6 ± 0.4 |
| Medium QC %CV | 2.8 | 1.9 | 4.3 | 4.0 | 3.6 | 3.3 | 2.4 | 2.0 |
| High QC (74.1 ng/mL) Mean +/- SD | 74.5 ± 1.8 | 74.5 ± 1.7 | 75.2 ± 2.2 | 77.1 ± 2.7 | 76.2 ± 2.5 | 71.6 ± 2.1 | 73.6 ± 1.8 | 73.7 ± 1.3 |
| High QC %CV | 2.3 | 2.0 | 2.9 | 3.5 | 3.3 | 3.0 | 2.5 | 1.7 |
| LoB ng/mL | 0.3 | 0.0 | 0.2 | 0.4 | * | N/A | 0.4 | 1.2 |
| LoD ng/mL | 0.7 | 0.5 | 0.6 | 0.6 | * | 0.7 | 0.7 | 1.7 |
| LoQ ng/mL | 5.0 | 4.0 | <6.0 | 4.0 | * | 2.1 | 4.0 | <4.0 |

* LoB, LoD and LoQ determined at site 3 on one instrument only.

Table 2: Galectin-3 reference interval: Log transformed upper 75th, 95th, and 97.5th percentile in all healthy individuals and stratified by gender and age, less than or equal or greater than 50 years of age.

|  | N | Mean (ng/mL) | Median (ng/mL) | 75th percentile (ng/mL) | 95th Percentile (ng/mL) | 97.5th percentile (ng/mL) |
| --- | --- | --- | --- | --- | --- | --- |
| All | 627 | 14.6 | 13.3 | 17.5 | 25.2 | 28.4 |
| Males | 335 | 14.8 | 13.7 | 17.8 | 25.9 | 33.3 |
| Females | 292 | 14.3 | 13.1 | 17.1 | 24.4 | 27.3 |
| <50 years | 457 | 13.7 | 12.5 | 16.3 | 22.9 | 25.5 |
| ≥50 years | 170 | 16.8 | 15.0 | 20.5 | 30.8 | 35.1 |

- Table 3: Biological variability of Galectin-3.

|  | Short term (within day) | Long term (5 days) | Combined Overall |
| --- | --- | --- | --- |
| Mean Galectin-3 ng/mL | 12.9 | 10.9 | 11.4 |
| Analytical variation, CVa% | 4.1 | 4.1 | 6.2 |
| Within subject biological variation, CVi% | 10.3 | 13.1 | 14.4 |
| Between subject biological variation, CVg% | 22.4 | 26.8 | 23.3 |
| Index of individuality, II | 0.50 | 0.50 | 0.67 |
| Reference Change Value, RCV | 30.7 | 38.0 | 43.5 |